## Summary
Evaporant name | Sm | Density (g/cm3) | 7.52
Atomic number | 62 | Z factor | 0.89
Atomic/molar mass (g/mol) | 150.4 | System | GM1
Physical properties |  | Operational details
Tm (oC) | 1072 | Ramp temp T1 (oC) | 300
Tv (oC) at 1e-8 Torr | 373 | Temp. change rate 1 (oC/ min) | 5 | (RT ~ Ti)
Tv (oC) at 1e-4 Torr | 573 | Ramp temp. T2 (oC) | 300
Chemical properties |  | Temp. change rate 2 (oC/ min) | 5 | (Ti ~ T')
Sensitivity to air | sensitive | Estimated Tg (oC) | 673
Sensitivity to moisture | sensitive | Temp. change rate 3 (oC/ min) | 5 | (T' ~ Tg)
Toxicity to human body | N/A | Maximum set point (oC) | 970
Flamability | N/A | Expected growth rate (A/s) | 0.01
Radioactivity | N/A | Source shape | dendritic pcs
Conditions to avoid | Incompatible materials (acids, air, water/moisture, oxidizing agents, acid chlorides, halogens) | Sealing / Handling | in Ar | in glovebox
Method of disposal |  | Cell type / Crucible volume (cc) | LT | 23 cc (LT/ST/DF)
|  | Cell position / Crucible material | cell 3 | Al2O3
Notes
Molten source
Evaporation category | Category I: Materials which sublime/can be evaporated in solid phase.
Notes | Use gloves when handling to prevent source contamination.
Past records | 0
Date | 2020-06-22 | Prepared by | Gordon Kentish

## Database
Material | Atomic # | Molar mass (g/mol) | Density (g/cm3) | z factor | Safety info summary | Sensitivity to air | Sensitivity to moisture | Toxicity to human body | Flamability | Radioactivity | Conditions to avoid | Tm (oC) | Tv @e-8 Torr (oC) | Tv @e-4 Torr (oC) | Estimated idle temp. Ti (oC) | Preferred Ti (exp. confirmed) | Estimated growth temp. Tg (oC) | Preferred Tg (exp. confirmed) GM1 | Preferred Tg (exp. confirmed) GM2 | Thermal evap. | Crucible materials | e-beam | Crucible liner materials | suitable cell | Preferred cell (exp. confirmed) | Evaporation history (Previous condition (Period, cell type/crucible material, Ti(oC)&Tg(oC) or power(%)&Ie(mA), flux (A/s)) | evaporation notes | Material Safety Data Sheet | Database completion | Database history (Any change on the values in the database as well as the date and reason of the change) | Current spec sheet GM1 | Current spec sheet GM2
Na | 11 | 22.98 | 0.968 | 4.8 | Danger: F(1), A(1, 1B) | Yes | Yes (potentially explosive) | On ingestion and contact with moisture on skin, eyes and mucous membranes to form sodium hydroxide leading to chemcal burns | Yes (approx. 290 degC ) or reaction with water. DO NOT USE WATER FIRE EXTINGUISHER | N/A | Exposure to any mositure. | 98 | 74 | 193 | 20 |  | 48 |  |  | Category III | Quartz | - | - | LT |  |  | Preheat gently to outgas. Metal reacts quickly in air | Added
Mg | 12 | 24.3 | 1.738 | 1.61 | Danger: F(1) (chip) / F(1, 2) (powder) | Readily forms oxidation layer | Yes | N/A | Yes in powder, shavings or molton form | N/A | Exposure to acids | 650 | 185 | 327 | 85 |  | 427 |  |  | Category I | Al2O3 | Good | Graphite, W | LT |  |  | https://www.lesker.com/newweb/deposition_materials/materialdepositionchart.cfm?pgid=0#s | Added
Al | 13 | 26.98 | 2.7 | 1.08 | Danger: F(1,2) | N/A | Emit flammable gases in contact with water | N/A | N/A | N/A | Exposure to moisture | 660 | 677 | 1010 | 517 |  | 610 |  |  | Category III | PBN, TiB2-BN | Excellent | Graphite | DF/e-beam |  |  | Alloys W/Mo/Ta. Flash evap or use BN crucible | Added
Si | 14 | 28.08 | 2.329 | 0.712 | N/A | Forms thin surface layer | N/A | Irritant on inhalation, ingestion and contact with skin/eyes | N/A | N/A | N/A | 1410 | 992 | 1337 | 832 |  | 1360 |  |  | Category III |  | Fair | Tantalum | HT |  |  | Alloys with W. Use heavy W boat. SiO produced.
K | 19 | 39.09 | 0.862 |  | Danger: F(1), A(1B) | Forms potassium peroxide | Yes Risk of Fire | Formation of KOH in contact with moisture on skin/eyes/mucous membranes. | Yes. DO NOT USE WATER FIRE EXTINGUISHER | N/A | Exposure to moisture | 64 | 22 | 124 | 20 |  | 14 |  |  | Category III | Al2O3, Quartz, PBN | - | - | LT |  |  | Metal reacts rapidly in air. Preheat gently to outgas. | Added
Ca | 20 | 40.08 | 1.54 | 2.62 | Danger: F(2) | Yes | Yes | N/A | Yes | N/A | Exposure to moisture | 839 | 272 | 459 | 172 |  | 559 |  |  | Category I | Al2O3, Quartz, PBN | Poor | - | LT |  |  | https://www.lesker.com/newweb/deposition_materials/deposition-materials-notes.cfm?pgid=ca1 | Added
Sc | 21 | 44.95 | 2.99 | 0.91 | Warning: F(2) | Passivation layer | N/A | N/A | Burns if ignited | N/A | N/A | 1541 | 714 | 1002 | 614 |  | 1102 |  |  | Category I | Al2O3 | Excellent | W, Mo | DF/e-beam |  |  | Alloys with Ta | Added
Ti | 22 | 47.87 | 4.5 | 0.628 | N/A | Yes | Yes | N/A | N/A | N/A | Exposure to Acids. Strong acids. Halogens. oxygen. Metals. Carbon dioxide (CO2). halocarbons. | 1660 | 1067 | 1453 | 967 | 1100 | 1553 | 1650 | 1340 | Category I | TiC, TiB2-BN, Ti-coated graphite (specially provided by DCA) | Excellent |  | HT/e-beam | HT | 03/02/2020, HT (2240-03) / DCA, Ti = 1100 degC, Tg = 1650 degC, ~0.02 A/s Crucible Material changed from Graphite to W 30/05/17, and then to Ti (DCA) 03/02/2020 |  | Added
V | 23 | 50.9415 | 5.96 | 0.53 | N/A | Readily form passivation layer. | N/A | N/A | N/A | N/A | N/A | 1890 | 1162 | 1547 | 1062 | 1300 | 1647 | 1800 | 1720 | Category I | Al2O3 | Excellent |  | HT/e-beam | HT/e-beam | - ~ 10/01/2020, HT (2240-02), W, Ti = 1300degC&Tg = 1700 degC, ~0.01 A/s - crucible was stuck (welded?) on the PBN cup at the bottom of the cell - 04/10/2019 ~, e-beam, W, | Annealing ex-situ in UHV at 170-200 C for 2 days is required to avoid huge outgassing during the first heating as experinenced in GM1. Not celar if this issue is due to contaminated V chunks purchased from 'American Elements' | Added | Completed
Cr | 24 | 51.996 | 7.2 | 0.305 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 1857 | 837 | 1157 | 737 | 750 | 1257 | 1280 | 1100 | Category I | vitreous carbon, Al2O3 | Good | Graphite, W, Al2O3 | HT | HT | 04/04/2019 ~, HT(9808-7)/Al2O3, Ti = 850 degC&Tg = 1280 degC, ~ 0.04 A/s | https://www.lesker.com/newweb/deposition_materials/deposition-materials-notes.cfm?pgid=cr1 | Added | Completed |  |  | https://www.dropbox.com/s/7zej6k70uowy37s/Cr.pdf?dl=0
Mn | 25 | 54.93 | 7.21 | 0.377 | N/A | N/A | N/A | Excessive exposure leads to toxicity. Risk with dust/fume inhalation | N/A | N/A | N/A | 1244 | 505 | 647 | 405 |  | 747 |  |  | Category I | Al2O3 | Good | W | ST |  |  |  | Added
Fe | 26 | 55.84 | 7.874 | 0.349 | Danger: F(1) | N/A | N/A | N/A | Flammable solids (H228, cat 1) Self-heating substances and mixtures (H251, cat 1) | N/A | N/A | 1535 | 858 | 1180 | 758 |  | 1280 |  |  | Category I | Al2O£ | Excellent | Fabmate | HT/e-beam |  |  | https://www.lesker.com/newweb/deposition_materials/deposition-materials-notes.cfm?pgid=fe1 | Added
Co | 27 | 58.93 | 8.86 | 0.343 | Danger: T(1), N(Ch 4) | Passivation layer | N/A | Respiratory sensitisation (H334, cat 1) Skin sensitisation (H317, cat 1) | N/A | N/A | N/A | 1495 | 850 | 1200 | 750 | 750 | 1300 | 1300 |  | Category I | Al2O3 | Excellent |  | HT/e-beam | HT | 30/01/2020, HT, Al2O3 Ti = 950 degC; Tg = 1300 degC, ~ 0.01 A/s | https://www.lesker.com/newweb/deposition_materials/deposition-materials-notes.cfm?pgid=co1 | Added
Ni | 28 | 58.693 | 8.912 | 0.331 | Danger: T(1), C(2) | N/A | N/A | - suspected of causing cancer (H317, cat. 1) - specific target organ toxicity by repeated exposure (H351, cat. 2) - may cause an allergic skin reaction (H372, cat. 1) | N/A | N/A | dust formation and ingestion or inhalation/contact with skin, eyes, clothing | 1453 | 927 | 1262 | 827 | 800 | 1362 | 1300 |  | Category I | Al2O3 | Excellent | Graphite, Al2O3 | HT/e-beam |  | >>> 04/10/2020 - 24/02/2020: e-beam, W, 4 - 7 % & 22 - 32 mA, ~ 0.03 A/s W crucible once broke due to rapid change in ebeam power, requires slow operation (< 1 %/min) >>> 24/02/2020 - 07/2020, DF (2240-11), Al2O3, no flux was seen in QCM >>> 07/2020 - : HT(2240-02), Al2O3 | https://www.lesker.com/newweb/deposition_materials/deposition-materials-notes.cfm?pgid=ni1 | Added | Completed
Cu | 29 | 63.54 | 8.96 | 0.437 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 1083 | 722 | 1017 | 562 |  | 1033 |  |  | Category III | Al2O3, Ta, Mo | Excellent | Graphite, Mo | DF/e-beam |  |  |  | Added
Zn | 30 | 65.38 | 7.14 | 0.514 | N/A | Passivation layer | N/A | Minor irritation | N/A | N/A | N/A | 420 | 123 | 247 | 20 |  | 370 |  |  | Category III | Al2O3, Quartz | Excellent | Fabmate, Graphite, W | LT |  |  | Evaporates well under wide range of conditions | Added
Ga | 31 | 69.72 | 5.91 |  | Warning: A(1) | N/A | N/A | Minor irritation. Wear gloves when handling. | N/A | N/A | N/A | 30 | 619 | 907 | 80 | 500 | 907 |  |  | Category II | Al2O3, Quartz | Good | Fabmate | ST |  |  | Alloys with W/Ta/Mo. Use E-beam gun. Low Melting Point materials not ideal for sputtering. | Added
Ge | 32 | 72.63 | 5.32 | 0.516 | N/A | Oxidises in air at 250 degC | N/A | Minor irritation | N/A | N/A | N/A | 937 | 812 | 1137 | 652 |  | 887 |  |  | Category III | Al2O3, Quartz | Excellent | Fabmate, Graphite | DF/e-beam |  |  |  | Added
As | 33 | 74.92 | 5.72 |  | Danger: T(3), C(1B), N(Ac 1, Ch 1) | N/A | Oxidises when exposed to moisture | Toxic if swallowed (H301) Toxic if inhaled (H331) May caused cancer (H350) | Flammable if heated. | N/A | Exposure to acids. | 817 | 104 | 210 | 4 |  | 310 |  |  | Category I | Al2O3 | Poor | Fabmate, PBN | LT |  |  | Sublimes rapidly at low temp. Not recommended for sputtering. | Added
Se | 34 | 78.97 | 4.81 | 0.864 | Danger: T(2 - 3), N(Ch 4) | N/A | N/A | Toxic if swallowed or inhaled. Organ damage through prolonged or extended exposure. Use protective wear | N/A | N/A | N/A | 217 | 63 | 164 | 20 |  | 167 |  | 150/500 | Category III | Al2O3, PBN | Poor | Fabmate, Mo | LT | Valved cracker cell |  |  | Added
Rb | 37 | 85.46 | 1.532 |  | Danger: T(1B), F(1) | Yes | Yes (dangerously) | Prolonged exposure risk. Risk of chemical burns if contact with moisture | Yes | N/A | Exposure to air or mositure. | 39 | -3 | 111 | -163 |  | -11 |  |  | Category III | Quartz | - | - | LT |  |  |  | Added
Sr | 38 | 87.6 | 2.64 | 1 | Danger: F(1) | sensitive | reactive produces flammable gas | N/A | Yes on reaction with water | N/A | Incompatible products (Acids, Oxidizing agents, water). Excess heat. | 769 | 239 | 403 | 139 | 180 | 503 | 350 | 350 | Category I | vitreous carbon, Mo, Al2O3* | Poor | - | LT | LT | -14/01/2020, LT (2240-07) / Al2O3, Ti = 180 degC, Tg = 300 - 420 degC, ~0.03 - 0.07 A/s -Al2O3 5 mm aperture added to cell with W spring to secure (04/06/2019) -Previous Ta appertures attempted resulted in chocking due to Sr buildup. |  | Added
Y | 39 | 88.9 | 4.472 | 0.835 | N/A | Passivation layer in bulk. Fine shavings can ignite in air | Reacts to form oxide. | Dangerous on inhalation. | Flammable if heated. | N/A | Excess heat. Strong acids. | 1522 | 830 | 1157 | 730 |  | 1257 |  |  | Category I | Al2O3 | Excellent | W | HT/e-beam |  |  | High Ta solubility | Added
Zr | 40 | 91.22 | 6.52 | 0.6 | N/A | N/A | N/A | Mild irritation. | Powder form. | N/A | N/A | 1852 | 1477 | 1987 | 1317 |  | 1802 |  |  | Category III | - | Excellent | W | e-beam |  |  | https://www.lesker.com/newweb/deposition_materials/deposition-materials-notes.cfm?pgid=zr1 | Added
Nb | 41 | 92.9 | 8.55 | 0.492 | N/A | N/A | N/A | N/A | N/A | N/A | In compatible products (Strong bases, Strong oxidizing agents, Halogens), attacks W source | 2468 | 1728 | 2287 | 1628 |  | 2387 |  |  | Category I | - | excellent | Fabmate | e-beam | e-beam | 18/06/20 Flux at 1.5nA for a 70min growth of sub monolayer coverage. GM2 e beam |  | Added
Mo | 42 | 95.9 | 10.2 | 0.257 | N/A | sensitive | N/A | N/A | N/A | N/A | Incompatible products (Acids, Oxidizing agents). Excess heat. | 2610 | 1592 | 2117 | 1492 |  | 2217 |  |  | Category I | - | excellent | Fabmate, graphite | e-beam |  |  | https://www.lesker.com/newweb/deposition_materials/deposition-materials-notes.cfm?pgid=mo1 | Added
Ru | 44 | 101.1 | 12.3 | 0.182 | N/A | N/A | N/A | N/A | N/A | N/A | excess heat | 2310 | 1780 | 2260 | 1620 |  | 2260 |  |  | Category III | - | Poor | - | e-beam | e-beam | 06/01/2017 ~, e-beam/W, 140 mA, ~ 0.06 A/s (QCM2) | https://www.lesker.com/newweb/deposition_materials/depositionmaterials_evaporationmaterials_1.cfm?pgid=ru1 | Added
Rh | 45 | 102.9 | 12.41 | 0.21 | Danger: F(1) | N/A | N/A | This product is or contains a component that is not classifiable as to carcinogenicity based on its IARC, ACGIH, NTP, or EPA classification. | Yes (powder) | N/A | Heat, flames, sparks. Incompatible materials (Strong oxidizing agents, Halogens, Chlorine trifluoride, Oxygen difluoride) | 1963 | 1277 | 1707 | 1177 |  | 1807 |  |  | Category I | ThO2, vitreous carbon | good | W, Fabmate (Kurt J Lesker) | HT | e-beam |  |  | Preliminary
Pd | 46 | 106.4 | 12.02 | 0.357 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 1554 | 842 | 1192 | 742 | 750 | 1292 | 1290 | 1100 | Category I | Al2O3 | Excellent | W, C, Graphite | HT/e-beam | HT | 19/12/2016 ~, HT(2240-01)/Al2O3, Ti = 850 degC, Tg = 1280 degC, ~ 0.016 A/s |  | Added | Completed
Ag | 47 | 107.8682 | 10.5 | 0.529 | Warning: N(Ac 1, Ch 1) | sensitive | N/A | N/A | N/A | N/A | exposure to air, (strong) acids, strong bases, ammoia, oxygen | 962 | 847 | 1105 | 687 |  | 912 |  |  | Category III | Al2O3, W, Mo | Excellent | W, Mo, Ta, C, Fabmate (Kurt J Lesker) | DF/e-beam |  | - 25/09/2019 ~ 03/02/2020, DF (9808-03), Al2O3, the crucible was broken due to rapid change in temperature across the melting tempearture - 24/02/2020 ~ , e-beam, W | https://www.lesker.com/newweb/deposition_materials/deposition-materials-notes.cfm?pgid=ag1 | Added | Completed |  | https://www.dropbox.com/s/1rfin4h2mltstny/Ag.pdf?dl=0
In | 49 | 114.82 | 7.31 | 0.841 | Warning: T(3 - 4), A(2) | N/A | N/A | Acute toxicity, Oral (Category 4), H302 Acute toxicity, Inhalation (Category 4), H332 Acute toxicity, Dermal (Category 4), H312 Skin irritation (Category 2), H315 Eye irritation (Category 2), H319 Specific target organ toxicity - single exposure (Category 3), H335 | N/A | N/A | Incompatible materials (Strong oxidizing agents, Sulphur compounds, Strong acids, Halogens, Acetonitrile, Tellurium, arsenic powder, phosphorous) | 157 | 487 | 742 | 207 |  | 742 |  |  | Category II | Al2O3, PBN, Graphite | Excellent | Graphite, Mo, Fabmate (Kurt J Lesker) | ST |  |  | https://www.lesker.com/newweb/deposition_materials/deposition-materials-notes.cfm?pgid=in1 | Preliminary
Sn | 50 | 118.69 | 7.28 | 0.724 | N/A | N/A | N/A | N/A | N/A | N/A | Incompatible materials (Acids. Strong oxidizing agents. Strong bases. Halogens. Metals.) | 232 | 682 | 997 | 282 |  | 997 |  |  | Category II | Al2O3, PBN | excellent | Ta, Fabmate (Kurt J Lesker) | DF/e-beam |  |  |  | Added
Sb | 51 | 121.75 | 6.68 | 0.768 | N/A | sensitive | N/A | N/A | N/A | N/A | Incompatible materials (Oxidizing agents, Acids, Hydrogen gas, halogen acids, Reducing agents) | 630 | 279 | 425 | 179 |  | 525 |  |  | Category I | PBN, Al2O3, C | poor |  | LT |  |  | -suggested valved cracker cell (createc evaporation guide) | Preliminary
Te | 52 | 127.6 | 6.25 | 0.9 | Warning: T(4), A(1), N(Ch 4) | N/A | N/A | H332 - Harmful if inhaled H317 - May cause an allergic skin reaction | N/A | N/A | Incompatible products (Strong oxidizing agents). Excess heat. Avoid dust formation. | 449 | 157 | 277 | -3 | 150 | 399 |  | 410 | Category III | PBN, Al2O3, Quartz | poor | Fabmate (Kurt J Lesker) | LT |  |  | -suggested valved cracker cell (createc evaporation guide) | Added
Cs | 55 | 132.905 | 1.88 | 1 | Danger: F(1), A(1B) | sensitive | reacts violently with water | Material is extremely destructive to tissue of the mucous membranes and upper respiratory tract, eyes, and skin, Cough, Shortness of breath, Headache, Nausea. H314 (Causes severe skin burns and eye damage.) | H260 (In contact with water releases flammable gases which may ignite spontaneously.). Not flammable in solid form. | N/A | Exposure to moisture. Incompatible materials (Do not store near acids., Phosphorus, Halogens, Alcohols, Oxygen, Reacts violently with water.) | 28 | -16 | 80 | -176 |  | -22 |  |  | Category III | Quartz, PBN | - | - | LT |  |  |  | Preliminary
Ba | 56 | 137.33 | 3.51 | 2.1 | Danger: F(2) | sensitive | reacts violently with water | Stomach/intestinal disorders, Nausea, Vomiting, Drowsiness, Dizziness, Gastrointestinal disturbance, Weakness, Tremors, Seizures. | H261 (In contact with water releases flammable gases.) | N/A | Exposure to moisture. In compatible materials (Oxidizing agents, Water, acids, Oxygen, Chlorinated solvents, Carbon dioxide (CO2), Halogens, Halogenated hydrocarbon, Alcohols, Sulphur compounds, Hydrogen sulfide gas) | 725 | 272 | 462 | 172 |  | 562 |  |  | Category I | W, Ta, Mo | fair | - | LT |  |  | Wets without alloying, reacts with ceramics. Not recommended for sputtering. | Added
La | 57 | 138.9055 | 6.15 | 0.92 | Danger: F(1) | sensitive | violently reactive produces flammable gas | N/A | Yes on reaction with water | N/A | Exposure to moisture. Exposure to incompatible materials (Strong acids, Strong oxidizing agents, Halogens, Phosphorus, Sulphur compounds) | 921 | 990 | 1388 | 971 | 1021 | 1388 | 1400 | 1400 | Category II | Al2O3, W* | Excellent | W, Ta | HT/e-beam | HT | - 10/01/2020 ~ , HT(2240-02)/Al2O3, crucible was broken due to rapid cooling - 24/02/2020 ~, HT(2240-02)/W, |  | Added
Ce | 58 | 140.12 | 6.77 | 1 | Danger: F(1) | sensitive | N/A | N/A | H228 (Flammable solids (Category 1)) | N/A | Heat, flames, sparks. Incompatible materials (Strong acids, Halogens, Strong oxidizing agents, Zinc) | 795 | 970 | 1380 | 845 |  | 1380 |  |  | Category II | Al2O3, vitreous C, BeO | good | - | HT |  |  |  | Added
Pr | 59 | 140.91 | 6.77 | 1 | Danger: F(1) | sensitive | reacts violently with water | N/A | H250 (Catches fire spontaneously if exposed to air.) | N/A | Air sensitive. Incompatible materials (Strong acids, Strong oxidizing agents, Halogens, water) | 931 | 800 | 1150 | 640 |  | 881 |  |  | Category III | - | good | - | HT |  |  | - | Preliminary
Nd | 60 | 144.24 | 7.01 | 1 | Warning: A(2), T(3) | sensitive | sensitive | N/A | N/A | N/A | N/A | 1021 | 731 | 1062 | 571 | 600 | 971 |  |  | Category III | Al2O3, W | Excellent | Ta | DF/e-beam |  | - 07/01/2020 ~ 24/02/2020, ST (2240-09), Al2O3, crucible was broken due to rapid temperature change across the melting temperature |  | Added | Completed
Sm | 62 | 150.4 | 7.52 | 0.89 | Danger: F(2), T(2) | sensitive | sensitive | N/A | N/A | N/A | Incompatible materials (acids, air, water/moisture, oxidizing agents, acid chlorides, halogens) | 1072 | 373 | 573 | 273 |  | 673 |  |  | Category I | Al2O3 | Good | - | ST |  |  |  | Added
Eu | 63 | 151.96 | 5.24 | 1 | Danger: F(1, 2) | sensitive | sensitive | N/A | H250 (Catches fire spontaneously if exposed to air.) H261 (In contact with water releases flammable gases, May form combustible dust concentrations in air) | N/A | Exposure to air. Incompatible products. Exposure to moist air or water. Incompatible products (Strong oxidizing agents. Strong acids. Halogens.) | 822 | 280 | 466 | 180 |  | 566 |  |  | Category I | Al2O3, Mo, Ta, PBN | Fair | - | LT |  |  |  | Added
Gd | 64 | 157.25 | 7.9 | 0.67 | Warning: F(3) | sensitive | reacts violently with water | N/A | H261 (In contact with water releases flammable gases.) | N/A | Exposure to mositure. Incompatible materials (Oxygen, Strong acids, Strong oxidizing agents) | 1313 | 760 | 1175 | 600 |  | 1263 |  |  | Category III | Al2O3 | Excellent |  | HT/e-beam |  |  |  | Added
Tb | 65 | 158.93 | 8.23 | 0.66 | Danger: F(1) | N/A | N/A | N/A | H228 (Flammable solid) | N/A | Heat, flames and sparks. Incompatible materials (Halogens, Strong acids) | 1356 | 800 | 1150 | 700 |  | 1250 |  |  | Category I | Al2O3 | Excellent | Graphite, Ta, Fabmate (Kurt J Lesker) | HT/e-beam |  |  | Wets without alloying. Not recommended for sputtering. | Preliminary
Dy | 66 | 162.5 | 8.55 | 0.6 | N/A | sensitive | sensitive | N/A | N/A | N/A | Incompatible materials (Acids, Oxidizing agents, Water, air) | 1412 | 625 | 900 | 525 |  | 1000 |  |  | Category I | Ta | good | - | ST |  |  |  | Preliminary
Ho | 67 | 164.93 | 8.8 | 0.58 | N/A | N/A | N/A | N/A | N/A | N/A | Incompatible materials (Strong oxidizing agents) | 1474 | 650 | 950 | 550 |  | 1050 |  |  | Category I | - | good | - | DF |  |  |  | Preliminary
Er | 68 | 167.26 | 9.07 | 0.74 | N/A | N/A | N/A | N/A | N/A | N/A | Incompatible materials (Strong acids, Strong oxidizing agents, Halogens) | 1497 | 650 | 930 | 550 |  | 1030 |  |  | Category I | PBN, Ta, W | good | Ta, W | ST |  |  |  | Preliminary
Tm | 69 | 168.93 | 9.32 | - | N/A | sensitive | sensitive | Skin & eye irritation | H228 (Flammable solid) H260 (In contact with water releases gases which may ignite spontaneously) | N/A | Incompatible materials (Acids, Oxidizing agents, Air, Halogens, Acid chlorides, Water/moisture) | 1545 | 461 | 680 | 361 |  | 780 |  |  | Category I | Al2O3 | good | - | ST |  |  |  | Preliminary
Yb | 70 | 173.04 | 6.96 | 1.13 | N/A | N/A | N/A | N/A | N/A | N/A | Incompatible materials (Strong oxidizing agents) | 819 | 247 | 417 | 147 |  | 517 |  |  | Category I | Ta | good | Ta | LT |  |  |  | Preliminary
Lu | 71 | 174.97 | 9.84 | - | N/A | sensitive | sensitive | Lantanons can cause delayed blood clotting leading to hemorrages. Exposure may also lead to sensitivity to heat, itching, increased awareness of odor and taste, and liver damage. | Powder flammable, contact with water releases extremely flammable gases (powder) | N/A | Incompatible materials (Oxidizing agents, Halogens, Air) | 1663 | 870 | 1376 | 770 |  | 1476 |  |  | Category I | Al2O3 | excellent | - | HT/e-beam |  |  |  | Preliminary
Hf | 72 | 178.49 | 13.31 | 0.36 | N/A | N/A | N/A | N/A | N/A | N/A | Excess heat. Incompatible materials (Halogens, halocarbons) | 2227 | 2160 | 3090 | 2000 |  | 2177 |  |  | Category III | - | good | - | e-beam |  |  |  | Added
Ta | 73 | 180.95 | 16.6 | 0.262 | Warning: F(2), A(2), T(3) | N/A | N/A | H315 (Causes skin irritation) H319 (Causes serious eye irritation) H335 (May cause respiratory irritation) | H228 (Flammable solid) | N/A | Excess heat. Incompatible materials (Acids, Strong oxidizing agents) | 2996 | 1960 | 2590 | 1860 |  | 2690 |  |  | Category I | - | excellent | Graphite, Fabmate (Kurt J Lesker) | e-beam |  | GM2 e beam. Have struggled to get a flux reading previously. | https://www.lesker.com/newweb/deposition_materials/deposition-materials-notes.cfm?pgid=ta1 | Preliminary
W | 74 | 183.85 | 19.3 | 0.163 | Warning: F(2) | N/A | N/A | N/A | H228 (Flammable solid) | N/A | Excess heat. Avoid dust formation. Keep away from open flames, hot surfaces and sources of ignition. Incompatible materials (Strong oxidizing agents. Alcohols. Halogens.) | 3410 | 2117 | 2757 | 2017 |  | 2857 |  |  | Category I | - | good | - | e-beam |  |  | https://www.lesker.com/newweb/deposition_materials/deposition-materials-notes.cfm?pgid=w1 | Preliminary
Re | 75 | 186.21 | 21.02 | 0.15 | N/A | N/A | N/A | N/A | N/A | N/A | Incompatible materials (Strong oxidizing agents) | 3180 | 1928 | 2570 | 1828 |  | 2670 |  |  | Category I | - | poor | - | e-beam |  |  |  | Preliminary
Os | 76 | 190.23 | 22.6 | - | Danger: F(2), A(1, 2), T(3) | sensitive | N/A | H315 (Causes skin irritation) H318 (Causes serious eye damage) H335 (May cause respiratory irritation) | H228 (Flammable solid) | N/A | Heat, flames and sparks. Incompatible materials (Oxygen, Halogens, Phosphorous) | 3033 | 2170 | 2760 | 2070 |  | 2860 |  |  | Category I | - | fair | - | e-beam |  |  |  | Preliminary
Ir | 77 | 192.22 | 22.5 | 0.129 | N/A | N/A | N/A | N/A | N/A | N/A | Excess heat | 2410 | 1850 | 2380 | 1690 |  | 2360 |  |  | Category III | - | Fair | - | e-beam |  |  | https://www.lesker.com/newweb/deposition_materials/deposition-materials-notes.cfm?pgid=ir1 | Added
Pt | 78 | 195.09 | 21.45 | 0.245 | Warning: F(2) | N/A | N/A | N/A | H228 (Flammable solid) | N/A | Excess heat. Avoid dust formation. Keep away from open flames, hot surfaces and sources of ignition. Incompatible materials (Strong oxidizing agents) | 1768 | 1292 | 1747 | 1132 |  | 1718 |  |  | Category III | C, BeO, ThO2 | Excellent | Graphite, Fabmate (Kurt J Lesker) | HT/e-beam |  |  | https://www.lesker.com/newweb/deposition_materials/deposition-materials-notes.cfm?pgid=pt1 | Added
Au | 79 | 196.97 | 19.3 | 0.381 | N/A | N/A | N/A | N/A | N/A | N/A | Incompatible materilas | 1063 | 807 | 1132 | 647 |  | 1013 |  |  | Category III | Al2O3, BN, PBN, vitreous C, W | excellent | Mo, Fabmate (Kurt J Lesker) | DF/e-beam |  |  | https://www.lesker.com/newweb/deposition_materials/deposition-materials-notes.cfm?pgid=au1 | Added
Pb | 82 | 207.2 | 11.34 | 1.13 | Danger: T(1A, 4), C(2), N(ac 1, ch 1) | N/A | N/A | H302+H332 (Harmful if swallowed or if inhaled.) H351 (Suspected of causing cancer.) H360FD (May damage fertility. May damage the unborn child.) H373 (May cause damage to organs through prolonged or repeated exposure) | N/A | N/A | Incompatible materials (Strong acids) | 327 | 342 | 497 | 182 |  | 277 |  |  | Category III | Al2O3, Quartz, Mo, PBN, Ta, W | excellent | Fabmate (Kurt J Lesker) | LT |  |  |  | Added
Bi | 83 | 208.98 | 9.8 | 0.79 | N/A | sensitive | N/A | Possible - Kidney disorders, Liver disorders, Dermatitis | N/A | N/A | Incompatible materials (Oxidizing agents, Halogens, acids) | 271 | 330 | 520 | 170 |  | 221 |  |  | Category III | Al2O3, vitreous C, BeO, Quartz | excellent | Graphite, Fabmate (Kurt J Lesker) | LT |  |  |  | Added
PdO | N/A | [13 empty] | 0 |  | 0 |  |  | 0
[15 empty] | 0 |  | 0 |  |  | 0
[15 empty] | 0 |  | 0 |  |  | 0
[15 empty] | 0 |  | 0 |  |  | 0
[15 empty] | 0 |  | 0 |  |  | 0
[15 empty] | 0 |  | 0 |  |  | 0
[15 empty] | 0 |  | 0 |  |  | 0
*Category I: Materials which sublime/can be evaporated in solid phase, set max SP at 100 oC below Tm.
*Category II: Materials which are in liquid phase during evaporation, set max SP at either 100 oC above Tv @1e-4 Torr or operatinal max tem. of the cell. Cool/heat the cell at 1oC/min rate while the cell temperature is within 100 oC from Tm. Take out the cell regularly to investigate the crucible/cell condition.
*Category III: Avoid using as much as possible. Otherwise, cool/heat the cell at 1oC/min rate while the cell temperature is within 100 oC from Tm. Take out the cell regularly to investigate the crucible/cell condition.